Contents

2. DEVICE DESCRIPTION & PRODUCT SPECIFICATION, INCLUDING VARIANTS & ACCESSORIES	2

2.1 Descriptive information for the device	2

# 2. DEVICE DESCRIPTION & PRODUCT SPECIFICATION, INCLUDING VARIANTS & ACCESSORIES

## 2.1 Descriptive information for the device

(a) Device Description

The -86°C Freezers described in this manual are intended for the storage of Frozen Red Blood Cells in 40% Glycerol at –65°C for up to 10 years; Frozen Rejuvenated Red Blood Cells at –65°C for up to 10 years; Fresh Frozen Plasma at –65°C for up to 10 years. These products are suitable for Vaccine and Pharmaceutical Storage when –50°C to –86°C is determined as an acceptable storage condition. These products should only be used by an authorized and adequately trained personnel.

The –86°C Freezers (refer to Models for the specific model series) described in this manual are intended for the storage of Frozen Red Blood Cells in 40% Glycerol at –65°C for up to 10 years; Frozen Rejuvenated Red Blood Cells at –65°C for up to 10 years; Fresh Frozen Plasma at –65°C for up to 10 years. These products are suitable for Vaccine and Pharmaceutical Storage when –50°C to –86°C is determined as an acceptable storage condition. These products should only be used by an authorized and adequately trained personnel. This unit is not intended for use in classified hazardous locations, nor to be used for the storage of flammable inventory.

Indications:

No Indications are defined in the Installation and Operation manual. The –86°C Freezers described in this manual are intended for the storage of Frozen Red Blood Cells in 40% Glycerol at –65°C for up to 10 years; Frozen Rejuvenated Red Blood Cells at –65°C for up to 10 years; Fresh Frozen Plasma at –65°C for up to 10 years. These products are suitable for Vaccine and Pharmaceutical Storage when –50°C to –86°C is determined as an acceptable storage condition. These products should only be used by authorized and adequately trained personnel. This unit is not intended for use in classified hazardous locations, nor to be used for the storage of flammable inventory.

Instructions for Use:

The –86°C Freezers (refer to Models for the specific model series) described in this manual are intended for the storage of Frozen Red Blood Cells in 40% Glycerol at –65°C for up to 10 years; Frozen Rejuvenated Red Blood Cells at –65°C for up to 10 years; Fresh Frozen Plasma at –65°C for up to 10 years. These products are suitable for Vaccine and Pharmaceutical Storage when –50°C to –86°C is determined as an acceptable storage condition. These products should only be used by an authorized and adequately trained personnel.

Table 1: Various Configurations or Variants of the device

| Model | Capacity Cu.(L) | Capacity Cryoboxes (2 in.) | Electrical (voltage, frequency, current) |
| --- | --- | --- | --- |
| TDE30086LA | 422 | 300 | 115V, 60 Hz, 1 Phase |
| TDE30086LD | 422 | 300 | 208-230V, 60 Hz, 1 Phase |
| TDE30086LV | 422 | 300 | 230V, 50 Hz, 1 Phase |
| TDE40086LA | 549 | 400 | 115V, 60 Hz, 1 Phase |
| TDE40086LD | 549 | 400 | 208-230V, 60 Hz, 1 Phase |
| TDE40086LV | 549 | 400 | 230V, 50 Hz, 1 Phase |
| TDE50086LA | 682 | 500 | 115V, 60 Hz, 1 Phase |
| TDE50086LD | 682 | 500 | 208-230V, 60 Hz, 1 Phase |
| TDE50086LV | 682 | 500 | 230V, 50 Hz, 1 Phase |
| TDE60086LA | 816 | 600 | 115V, 60 Hz, 1 Phase |
| TDE60086LD | 816 | 600 | 208-230V, 60 Hz, 1 Phase |
| TDE60086LV | 816 | 600 | 230V, 50 Hz, 1 Phase |

• 2.2 Product specification**\n

-86 Ultra Low temperature Freezer is designed for efficiency. As a result, the series delivers a foundational experience with the reliability and performance needed to protect users work—and the sustainability to meet global compliance. The devices are available in different models, with their features, dimensions, and performance attributes summarized in the table below.

Table 2: Product specification

| Model | Temp range (°C) | Capacity (L) | Refrigerant | Connectivity | Power cord specifications | Electrical (voltage, frequency, current) | Product capacity | Interior dimensions D x W x H (cm) | Exterior dimensions D x W x H (cm) | Shipping weight (kg) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| TDE30086LA | -50°C to -86°C | 422 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | NEMA 5-20 / IEC Cords, 10 ft (3.05 m) | 115V, 60 Hz, 1 Phase | 300 Standard 2" Boxes | 45.2 x 71.9 x 130.0 | 58.4 x 97.8 x 198.1 | 323 |
| TDE30086LD | -50°C to -86°C | 422 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | NEMA 6-15P / IEC Cords, 10 ft (3.05 m) | 208-230V, 60 Hz, 1 Phase | 300 Standard 2" Boxes | 45.2 x 71.9 x 130.0 | 58.4 x 97.8 x 198.1 | 323 |
| TDE30086LV | -50°C to -86°C | 422 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | Country Dependant plug / IEC Cords, 10 ft (3.05 m) | 230V, 50 Hz, 1 Phase | 300 Standard 2" Boxes | 45.2 x 71.9 x 130.0 | 69.3 x 97.8 x 198.1 | 323 |
| TDE40086LA | -50°C to -86°C | 549 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | NEMA 5-20 / IEC Cords, 10 ft (3.05 m) | 115V, 60 Hz, 1 Phase | 400 Standard 2" Boxes | 58.7 x 71.9 x 130.0 | 71.9 x 97.8 x 198.1 | 312 |
| TDE40086LD | -50°C to -86°C | 549 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | NEMA 6-15P / IEC Cords, 10 ft (3.05 m) | 208-230V, 60 Hz, 1 Phase | 400 Standard 2" Boxes | 58.7 x 71.9 x 130.0 | 71.9 x 97.8 x 198.1 | 312 |
| TDE40086LV | -50°C to -86°C | 549 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | Country Dependant plug / IEC Cords, 10 ft (3.05 m) | 230V, 50 Hz, 1 Phase | 400 Standard 2" Boxes | 58.7 x 71.9 x 130.0 | 82.6 x 97.8 x 198.1 | 312 |
| TDE50086LA | -50°C to -86°C | 682 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | NEMA 5-20 / IEC Cords, 10 ft (3.05 m) | 115V, 60 Hz, 1 Phase | 500 Standard 2" Boxes | 73.2 x 71.9 x 130.0 | 86.4 x 97.8 x 198.1 | 333 |
| TDE50086LD | -50°C to -86°C | 682 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | NEMA 6-15P / IEC Cords, 10 ft (3.05 m) | 208-230V, 60 Hz, 1 Phase | 500 Standard 2" Boxes | 73.2 x 71.9 x 130.0 | 86.4 x 97.8 x 198.1 | 333 |
| TDE50086LV | -50°C to -86°C | 682 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | Country Dependant plug / IEC Cords, 10 ft (3.05 m) | 230V, 50 Hz, 1 Phase | 500 Standard 2" Boxes | 73.2 x 71.9 x 130.0 | 97.0 x 97.8 x 198.1 | 333 |
| TDE60086LA | -50°C to -86°C | 816 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | NEMA 5-20 / IEC Cords, 10 ft (3.05 m) | 115V, 60 Hz, 1 Phase | 600 Standard 2" Boxes | 87.4 x 71.9 x 130.0 | 100.6 x 97.8 x 198.1 | 372 |
| TDE60086LD | -50°C to -86°C | 816 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | NEMA 6-15P / IEC Cords, 10 ft (3.05 m) | 208-230V, 60 Hz, 1 Phase | 600 Standard 2" Boxes | 87.4 x 71.9 x 130.0 | 100.6 x 97.8 x 198.1 | 372 |
| TDE60086LV | -50°C to -86°C | 816 | R290 / R170+R290 Mix | RS485/4-20mA output/Dry Contacts | Country Dependant plug / IEC Cords, 10 ft (3.05 m) | 230V, 50 Hz, 1 Phase | 600 Standard 2" Boxes | 87.4 x 71.9 x 130.0 | 111.3 x 97.8 x 198.1 | 372 |